## Sheet1
Table S2: Metadata of LA-MC-ICPMS U-Pb and Lu-Hf data used in this study
Sample Preparation
Sample type/mineral | Detrital zircons
Sample preparation | Traditional heavy liquid and magnetic techniques, 1 inch resin mount, 1 mm polish to finish
Imaging | Gatan Mono CL 4+ spectrometer attached to FEI Quanta 450 field emission scanning Electron Microscope, 13.4 mm working distance
LA-ICPMS U-Pb analyses
Laboratory name | Department of Earth Sciences, University of Hong Kong
Laser ablation system
Make, Model and type | Resonetics LLC USA, RESOlution M-50, ArF excimer
Ablation cell and volume | Two-volume laser-ablation cell (Laurin Technic, Australia), effective volume ∼1–2 cm3
Laser wavelength (nm) | 193 nm
Pulse width (ns) | ~20 ns
Fluence (J cm-2) | 5 J cm-2
Repetition rate (Hz) | 6 Hz
Ablation duration (s) | 45 s
Ablation pit depth / ablation rate | not available
Spot diameter (mm) nominal/actual | 0.03 mm
Sampling mode / pattern | spot
Carrier gas | 100% He in the cell, N2 were introduced into the make-up gas via a T-connector
Cell carrier gas flow (l min-1) | 500 ml/min He + 2 ml/min N2, Ar make-up gas.
ICP-MS Instrument
Make, Model and type | Nu Instruments, Nu Plasma HR, MC-ICPMS
Sample introduction | Ablation aerosol
RF power (W) | 1300 W
Make-up gas flow (l min-1) | 400 ml/min Ar
Detection system | Mixed Faraday-multiple ion counting array
Masses measured | 172–179, 204, 206, 207, 238
Integration time per peak/dwell times (ms); quadrupole settling time between mass jumps | not available
Total integration time per output data point (s) | not available
‘Sensitivity’ as useful yield (%, element) | not available for quadrupole ICPMS
IC Dead time (ns) | not available
Data Processing
Gas blank | 30 s on-peak zero subtracted
Calibration strategy | 91500 used as calibration standard, GJ-1 used as quality control for validation
Reference Material info | 91500 (Wiedenbeck et al. 1995)
GJ-1 (Jackson et al. 2004)
Data processing package used / Correction for LIEF | Commercial software ICPMSDataCal was intially developed by Prof. Yong-sheng LIU of GUG. LIEF correction assumes reference material and samples behave identically, namely standard-samples-standard bracketing.
Mass discrimination | not available
Common-Pb correction, composition and uncertainty | No common-Pb correction applied to the data.
Uncertainty level and propagation | Ages are quoted at 2s absolute, propagation is processed by the ICPMSDataCal software. Reproducibility and age uncertainty of reference material are propagated.
Quality control / Validation | GJ-1 – Weighted ave 206Pb/238U age = 602.3 ± 1 (2σ, MSWD = 0.6, n = 25)
Systematic uncertainty for propagation is 1.5% (1s).
Other information | Sample line of 3.5 m from ablation cell to torch and a 30 s washout time after the laser stopped firing.
LA-MC-ICPMS Lu-Hf analyses
Laboratory and Sample Preparation
Laboratory name | State Key Laboratory of Geological Processes and Mineral Resources, China University of Geosciences (Wuhan)
Laser ablation system
Make, Model and type | Lambda Physik, Gttingen, Germany, GeoLas2005 laser ablation system
Ablation cell and volume | HelEx II active 2-volume ablation cell, ~5 cm3 for small cell
Laser wavelength (nm) | 193 nm
Pulse width (ns) | 5 ns
Fluence (J cm-2) | 5.3 J cm-2
Repetition rate (Hz) | 6 Hz
Ablation duration (s) | 40 s
Ablation pit depth / ablation rate | about 15μm
Spot diameter (mm) nominal/actual | 44 μm
Sampling mode / pattern | Static spot ablation
Carrier gas | 100% He in the cell, N2 were introduced into the make-up gas via a T-connector
Cell carrier gas flow (l min-1) | 0.85 L/min
MC-ICP-MS Instrument
Make, Model and type | Thermo Fisher Scientific, Germany, Neptune Plus MC-ICP-MS
Sample introduction | dry aerosol by laser ablation
RF power | 1270 W
Make-up gas flow (l min-1) | ca. 0.3-1.31 L/min
Detection system | Faraday cups
Masses measured | 171Yb, 173Yb, 174(Hf+Yb), 175Lu,176(Hf+Yb+Lu), 177Hf, 179Hf, 180(Hf+W), 182W
Integration time per peak/dwell times (ms); quadrupole settling time between mass jumps | 1s
Total integration time per output data point (s) | 70s
‘Sensitivity’ as useful yield (%, element) | not available
IC Dead time (ns) | not available
Data Processing
Gas blank | 20 s on-peak zero subtracted
Calibration strategy | 91500 used as calibration standard, GJ-1 used as quality control for validation
Reference Material info | 91500 (Wu et al. 2006)
GJ-1 (Yuan et al. 2008)
Data processing package used / Correction for LIEF | Commercial software ICPMSDataCal was intially developed by Prof. Yong-sheng LIU of GUG. LIEF correction assumes reference material and samples behave identically, namely standard-samples-standard bracketing.
Mass discrimination | internal calibration by 179Hf/177Hf and 173Yb/171Yb
Uncertainty level and propagation | Ages are quoted at 2s absolute, propagation is processed by the ICPMSDataCal software. Reproducibility and age uncertainty of reference material are propagated.
Quality control / Validation | 91500 – Weighted ave 177Hf/176Hf = 0.282309 ± 0.000009 (2s, MSWD = 0.9, n = 6). GJ-1 – Weighted ave 177Hf/176Hf = 0.282013 ± 0.000006 (2s, MSWD = 1.1, n = 13)
Systematic uncertainty for propagation is 0.0015% (1s).
Other information | ca. 2m sample line from ablation cell to torch and washout time was ca. 1s to 1% of peak signal.